Quantitative analysis of image quality for acceptance and commissioning of an MRI simulator with a semiautomatic method

DOI：10.12072/ncdc.imp.db2904.2023

CSTR：11738.11.NCDC.IMP.DB2904.2023

Share type:Apply to Access

Category:Others

Datasets description

This paper establishes a set of semi-automatic MRI acceptance debugging method, which can quantitatively analyze the acceptance modulation image quality of semi-automatic MRI, and can be used as the evaluation and reference standard of imaging simulator performance, which ensures the positioning accuracy and image quality of MRI simulation positioning system in clinical use, and is conducive to improving the accuracy and curative effect of treatment.

Data source description

Magnetic resonance imaging (MRI) simulation is different from diagnostic MRI in purpose, technical requirements and implementation. We propose a semi-automatic image receiving and debugging method for pulse sequences of scanners, RF coils and MRI simulators. ACR MRI certification model was used for image quality analysis with seven parameters. Standard ACR sequences with split head coils were used to check the basic performance of the scanner. The performance of analog RF coils was measured and compared with different clinical diagnostic coils using standard sequences. We use the simulated sequence with analog coils to test the image quality and the advanced performance of the scanner. The code and program are developed for semi-automatic image quality analysis. When using standard ACR sequences Split coil, image quality through all ACR recommendations. Compared with the analog RF coil, the image intensity uniformity of the analog RF coil is reduced by about 34%. The other six image quality parameters are acceptable. These two image quality parameters can be improved to more Through the built-in strength calibration method more than 85%. In the simulation sequence test, the contrast resolution is sensitive to FOV and matrix settings. Geometric distortion of simulated sequences, such as T1 weighted and T2 weighted images The isocenter was well controlled, and the range 10 cm away from the center was 1% (2 mm). We developed a semi-automatic image quality analysis method for quantitative image evaluation and debugging of MRI simulator. The baseline performance routine QA of analog RF coil and pulse train has been established. C í g ò ngzh è n ch é ngxi à ng (M

Data processing method

Standard ACR sequences with split head coils were used to check the basic performance of the scanner. The performance of analog RF coils was measured and compared with different clinical diagnostic coils using standard sequences. The simulated sequence with simulated coil is used to test the image quality and the advanced performance of the scanner. The code and program are developed for semi-automatic image quality analysis. When standard ACR sequences are used, the image quality passes all ACR recommendations.

Data quality description

Good data quality

Base information

collect time:2017/07/01-2019/12/31

collect place:

data size:740.5 KiB

data format:

Coordinate system: 无

Keywords

Theme:Semi automatic, MRI, image quality
	Subject:
	Time:2017 2018 2019
	Place:Gansu,  Gansu Province

Space scope

|  | north： |  |
| --- | --- | --- |
| west： |  | east： |
|  | south： |  |

Service information

contacts:Min Yufang

phone:0931-4967596

mailbox:ncdc@lzb.ac.cn

Contact information

contributors:XIAO Guoqing

author:XIAO Guoqing

manager:XIAO Guoqing

Project information

number:2017YFC0107500，
      name:Research and implementation of a new conformal intensity modulation technique for multi-particle biology guided by multi-modes

Citations and annotations

In order to protect the rights and interests of the platform's scientific and technological resources, expand the services of the platform center, and enhance the application potential of scientific and technological resources,resource users are requested to mark and quote the research results (including published papers, treatises, data products, unpublished research reports, data products, etc.) generated by the use of resources in the following manner.

For the achievements published in English, please refer to the following specifications：

The dataset is provided by National Cryosphere Desert Data Center(https://www.ncdc.ac.cn).

Data reference:

XIAO Guoqing. Quantitative analysis of image quality for acceptance and commissioning of an MRI simulator with a semiautomatic method. National Cryosphere Desert Data Center(http://www.ncdc.ac.cn), 2023. https://cstr.cn/CSTR:11738.11.NCDC.IMP.DB2904.2023.